{"text_format":"3","workbook":"xls","n_sheets":1,"sheets":["Solvent Compatibility"],"not_rendered":{"external_links":1,"defined_names":2},"read":"whole","date_system":"1900","formulas":"cached_only"}

{"sheet":"Solvent Compatibility","index":0,"n_rows":148,"n_cols":20,"rows_kept":148,"merged":[[16,1,16,4],[16,7,16,10],[16,12,16,13],[16,14,16,15]],"flags":["formulas_cached_only"]}
[]
["Solvents listed below may be used in the machines listed where:"]
["Y","indicates \"yes\" no restriction"]
["O","indicates \"only with option\" e.g. HCl resistance, or inert gas purge for di-ethyl ether - see notes column for details"]
["E ","EZ-2 Elite only"]
["-","not tested / no data"]
["N","idicates \"no\" this solvent may not be used"]
["< X %","indicates only where concentration is below X%"]
["2007-04-01","indicates only on systems built after this date, and not on systems built before"]
[]
["SampleGenie seal data is only a guide, and only relates to seal resistance to a particular solvent.   All seals will be effected by solvent over time - nothing is 100% resistant.  "]
["When chosing SampleGenie seals customers must also consider the effects of leachables, and the effect of their dissolved sample on the seal - if in doubt, chose the Perfluoro"]
["If the solvent is not listed, or you have any doubts about compatibility, please contact Genevac R&D for approval "]
["For Mega system compatibilty, please contact Genevac"]
[]
["2015-06-30","Last Updated"]
[null,"Solvent",{"$nulls":5},"HT series 2",{"$nulls":3},"S3-HT","Rocket",null,"SampleGenie",null,"Manuals"]
["Class","Systematic name","old name","CAS","Abbr","miVac","EZ-2","HT-4X","DD-4X","HT-12","HT-24",null,"Chemi","4D","Viton","Perfluoro",null,"Added","Notes","Authorised"]
["Acids","1,1,1,3,3,3-Fluoro-2-propanol","hexafluoroisopropanol",null,"HFIP","N","O","O","N","O","N","O","N","N","N","Y","N","2006-06-15","HCl resistant systems only, not in CVP","RD"]
[null,"Acetic acid",{"$nulls":3},"Y","Y","Y","Y","Y","Y","Y","Y","Y","N","Y","Y"]
[null,"Acetic anhydride",{"$nulls":3},"N","Y","Y","Y","Y","Y","Y","Y","Y","N","Y","Y","2008-04-02",null,"RD"]
[null,"Butanoic acid","butyric acid",null,null,"N","O","Y","Y","Y","Y","Y","Y","Y","N","Y","N","2013-10-24","EZ-2 Plus / Elite and HTs only.  Not miVac or EZ-2 std / Envi","RD"]
[null,"2-Ethylhexanoic acid","Butylethylacetic aicd",null,"EHA","N","E","Y","Y","Y","Y","Y","N","N","N","Y","N","2011-07-26","EZ-2 Elite only, or HT series.  May need high temperature","RD"]
[null,"Formic acid",null,null,"HCOOH","Y","Y","Y","Y","Y","Y","Y","Y","Y","N","Y","Y",null,"Not compatible with Cole pump unless CPC option used"]
[null,"Heptafluorobutyric acid","Perfluorobutanoic acid\nPerfluorobutyric acid",null,"HFBA","N","O","O","N","O","N","O","N","N","N","Y","N","2015-03-11","Requires HCl resistance option.  You may use up to 0.1% (but no more) in standard Genevac systems without a special option.  No HFBA can be used in any miVac","AW"]
[null,"Hydrochloric acid",null,null,"HCl","N","O","O","N","O","N","O","N","N","N","Y","N","2010-10-22","Requires HCl resistance option.  You may use up to 0.1% (but no more) in standard Genevac systems without a special option.  No HCl can be used in any miVac","RD"]
[null,"Hydrobromic acid",null,null,"HBr","N","O","O","N","O","N","O","N","N","Y","Y","Y",null,"HCl resistant systems only"]
[null,"Hydrofluoric acid",null,null,"HF","N","N","N","N","N","N","N","N","N","N","N","N"]
[null,"2-Methylpropanoic acid","Isobutyric acid",null,null,"N","O","Y","Y","Y","Y","Y","Y","Y","N","Y","N","2013-06-12","EZ-2 Plus / Elite and HTs only.  Not miVac or EZ-2 std / Envi"]
[null,"Nitric acid",null,null,"HNO3","N","O","N","N","N","N","N","N","N","N","Y","Y","2009-09-01","Only in EZ-2 Plus mark 3 with HCl option, no other systems, not EZ-2 standard, not EZ-2 Envi, not EZ-2 Elite","RD/RLS"]
[null,"Nitro-hydrochloric acid","Aqua Regia",null,null,"N","N","N","N","N","N","N","N","N","N","N","N","2013-04-10","Super acid - much stronger than HNO3","RD"]
[null,"Oxalyl chloride",null,null,"C2Cl2O2","N","O","O","N","O","N","O","N","N","N","Y","N","2006-08-25","HCl resistant systems only, not in CVP","RD"]
[null,"Pentafluoropropanoic acid",null,null,"PFPA","N","Y","Y","N","Y","Y","Y","Y","N","N","Y","N","2011-09-29","HCl resistant system advised","RD"]
[null,"Peroxyacetic acid",{"$nulls":3},"N","N","N","N","N","N","N","N","N","N","N","N","2007-03-20","Solvents are unstable to heat and may explode","RD"]
[null,"Peroxyformic acid","Performic acid",null,null,"N","N","N","N","N","N","N","N","N","N","N","N","2007-03-20","Solvents are unstable to heat and may explode","RD"]
[null,"Propanoic acid",{"$nulls":3},"N","O","Y","Y","Y","Y","Y","N","N","N","Y","N","2008-11-28","Amended to clarify.  EZ-2+/Elite only","RD"]
[null,"n-Pentyl propionate","Amyl propionate",null,null,"N","O","Y","Y","Y","Y","Y","N","N","N","Y","N","2008-11-28","EZ-2 Elite only, or HT series","RD"]
[null,"Metaphosphoric acid",null,null,"HO3P","N","*","N","N","N","N","N","N","N","N","Y","N","2006-08-21","* <0.5% approved with  HCl option for EZ-2 but NOT Elite model","RD/RLS"]
[null,"methane-sulfonic acid",{"$nulls":3},"N","N","N","N","N","N","N","N","N","N","N","N","2017-07-17","Boiling Point at atmosphere is +217°C, whereas the BP under vaccum at 13mbar is +167°C - cannot be evaporated in Genevac products"]
[null,"Phosphoric acid",null,null,"H3P04","N","*","N","N","N","N","N","N","N","Y ","Y","N","2006-08-21","* <0.5% approved with  HCl option for EZ-2 but NOT Elite model","RD/RLS"]
[null,"Perchloric acid",null,null,"PClO4","N","N","N","N","N","N","N","N","N","N","N","N","2006-08-18","too unstable, fire and explosion risk, also dissolves metals"]
[null,"Phosgene","dichloromethanone",null,"COCl2","N","O","O","N","O","N","O","N","N","N","Y","N","2012-05-24","HCl resistant systems only","RD"]
[null,"Phosphorus oxychloride",null,null,"PClO3","N","N","N","N","N","N","N","N","N","N","N","N","2009-03-18","unstable - dissolves most metals","RD"]
[null,"Sulphuric acid",null,null,"H2SO4","N","N","N","N","N","N","N","N","N","N","N","N","2006-08-18","too strong, dissolves metals too well, evolves Hydrogen"]
[null,"Thionyl chloride",{"$nulls":3},"N","O","O","N","O","N","O","N","N","N","Y","Y",null,"creates HCl, needs HCl option"]
[null,"Trifluoroacetic acid",null,null,"TFA","<0.5%","Y","Y","Y","Y","Y","Y",{"v":0.005,"d":"0.50%"},"<0.5%","N","Y","Y",null,"miVac conc corrected to 0.5% 20-Dec-11\nFurther testing required on Rocket 19-Aug-14","RD\nAW"]
[null,"Trifluoroacetic anhydride",{"$nulls":3},"N","O","O","N","O","N","O","N","N","N","Y","N","2011-03-04","evolves hydrogen in presence of water - needs care and customer must make own risk assesment","RD"]
[null,"Trichloroacetic acid",null,null,"TCA","<0.5%","O","O","N","O","N","O","N","N","N","Y","Y","2009-08-25","HCl option, very high BP therefore HCl resistant HTs / Elite only"]
[null,"2,2,2-Trifluoroethanol",null,null,"TFE","N","Y","Y","Y","Y","Y","Y","Y","N","N","Y","Y","2007-09-20","Treat as TFA"]
[null,"Trifluoromethanesulfonic acid",{"$nulls":3},"N","N","N","N","N","N","N","N","N","N","N","N","2010-02-24","Super acid - much stronger than HNO3","RD"]
[null,"Tris(hydroxymethyl)aminomethane hydrochloride","TRIShydrochloride ",null,"Tris HCl","N"]
[]
[]
["Bases","Ammonium acetate",null,null,"CH3COONH4","Y","Y","Y","Y","Y","Y","Y","Y","Y","N","Y","N","2006-08-04","Dissociates to form ammonia and acetic acid","RD"]
[null,"Ammonium bicarbonate",null,null,"(NH3)2CO3","Y","Y","Y","Y","Y","Y","Y","Y","Y","Y","Y","N"]
[null,"Ammonium hydroxide",null,null,"NH3OH","Y","Y","Y","Y","Y","Y","Y","Y","Y","N","Y","Y"]
[null,"Azacyclohexane","Piperidine",null,null,"2007-04-01","Y","Y","Y","Y","-","Y","Y","Y","N","Y","N","2007-04-03","May attack seals on miVac","RD"]
[null,"Diamine","Hydrazine",null,"H2N-NH2","N","N","N","N","N","N","N","N","N","N","N","N","2007-12-11","Explosive in presence of light sources, metals and shock!","RD"]
[null,"Dimethylaminoethanol",null,null,"DMAE","N","O","Y","Y","Y","Y","Y","Y","Y","N","Y","N","2009-04-15","EZ-2 plus / Elite only. Use odour reduction.","RD"]
[null,"Diisopropylethylamine",{"$nulls":3},"N","Y","Y","Y","Y","Y","Y","Y","Y","N","Y","N","2010-02-16",null,"RD"]
[null,"Methylamine (aqueous)",null,null,"CH5N(aq)","N","O","O","N","O","N","O","O","N","N","Y","N","2013-07-23","Must only be used with Inert Gas Purge","RD / RLS"]
[null,"Pyridine","Azabenzene, Azine",null,"C5H5N","N","O","Y","Y","Y","Y","Y","Y","Y","N","Y","Y","2012-05-23","EZ-2 plus / Elite only","RD"]
[null,"Sodium hydroxide (aqueous)",null,null,"NaOH","<2M","<2M","<2M","<2M","<2M","<2M","<2M","<2M","<2M","N","Y","Y","2008-11-11","Solution in water, <2 molar","RD"]
[null,"Triethylamine",{"$nulls":3},"<5%","Y","Y","Y","Y","Y","Y","Y","Y","N","Y","Y",null,null,"RD"]
[]
[]
[]
["Solvents","Acetonitrile",null,null,"MeCN","Y","Y","Y","Y","Y","Y","Y","Y","Y","N","Y","Y",null,null,"RD"]
[null,"Benzene",{"$nulls":3},"N","N","N","N","N","N","N","N","N","N","N","N","2007-01-31","Attacks polypropylene","RD"]
[null,"Bromine",null,null,"Br2","N","N","N","N","N","N","N","N","N","N","N","N","2010-02-26","Incompatible with most metals and organic materials","RD"]
[null,"1,3-Butanediol","Butylene Glycol",null,null,"N","N","Y","Y","Y","Y","Y","N","N","N","Y","N","2010-07-14","Very high BP(207), may need SG temp >40°C","RD"]
[null,"1,4-Butanediol","Butylene Glycol",null,null,"N","N","N","N","N","N","N","N","N","N","Y","N","2014-06-19","BP228","AW"]
[null,"2,3-Butanediol","Butylene Glycol",null,null,"N","E","Y","Y","Y","Y","Y","N","N","N","Y","N","2014-06-19","BP185","AW"]
[null,"Butan-1-ol",{"$nulls":3},"O","O","Y","Y","Y","Y","Y","Y","Y","Y","Y","Y",null,"EZ-2+ or Elite only / miVac Quattro Pump or scroll"]
[null,"Butan-2-ol",{"$nulls":3},"Y","Y","Y","Y","Y","Y","Y","Y","Y","Y","Y","Y"]
[null,"Butanone","methyl ethyl ketone",null,"MEK","Y","Y","Y","Y","Y","Y","Y","Y","Y","N","Y","Y"]
[null,"2-Butoxyethanol","ethylene glycol monobutyl ether",null,null,"N","E","Y","Y","Y","Y","Y","N","N","N","Y","N","2012-11-13",null,"RD"]
[null,"Butyl acetate",{"$nulls":3},"Apr-07&O","O","Y","Y","Y","Y","Y","Y","Y","N","Y","Y","2007-03-14","EZ-2+ or Elite only / miVac Quattro Pump or scroll","RD"]
[null,"Carbon disulphide","Carbon disulfide",null,"CS2","N","O","O","N","O","O","O","O","N","N","Y","N","2011-07-18","Must only be used with Inert Gas Purge","RD"]
[null,"Monochlorobenzine",{"$nulls":3},"N","O","Y","Y","Y","Y","Y","Y","Y","Y","Y","Y","2009-11-23","EZ-2 plus / Elite only"]
[null,"1-Chlorobutane","1-Butyl Chloride",null,null,"Y","Y","Y","Y","Y","Y","Y","Y","Y","Y","Y","N","2010-04-09",null,"RD"]
[null,"2-Chlorobutane","2-Butyl Chloride",null,null,"Y","Y","Y","Y","Y","Y","Y","Y","Y","Y","Y","N","2012-05-21",null,"RD"]
[null,"tert-Butyl chloride","t-Chlorobutane",null,null,"Y","Y","Y","Y","Y","Y","Y","Y","Y","Y","Y","N","2012-05-21",null,"RD"]
[null,"Chloroform",{"$nulls":3},"Y","Y","Y","Y","Y","Y","Y","Y","Y","Y","Y","Y"]
[null,"Cyanogen bromide",{"$nulls":3},"N","N","N","N","N","N","N","N","N","N","N","N","2011-03-25","Reacts with metals and with moisture","RD"]
[null,"Cyclohexane",{"$nulls":3},"Y","Y","Y","Y","Y","Y","Y","Y","Y","Y","Y","Y","2008-04-02",null,"RD"]
[null,"Cyclohexanone",{"$nulls":3},"N","O","Y","Y","Y","Y","Y","Y","Y","N ","Y","N","2012-03-26","EZ-2 plus / Elite only","RD"]
[null,"Cyclohexylmethanol",{"$nulls":3},"N","N","Y","Y","Y","Y","Y","Y","Y","Y","Y","N","2008-04-02",null,"RD"]
[null,"1-Decanol",{"$nulls":3},"N","N","N","N","N","N","N","N","N","N","N","N","2013-09-25","BP 231°C is too high","RD"]
[null,"1,2-Dibromoethane","ethylene dibromide",null,null,"Apr-07&O","O","Y","Y","Y","Y","Y","Y","Y","N","Y","N","2010-03-30","EZ-2 plus or Elite only / miVac Quattro Pump or scroll","RD"]
[null,"Dibromomethane",{"$nulls":3},"2007-04-01","Y","Y","Y","Y","Y","Y","Y","Y","Y","Y","N","2010-03-30",null,"RD"]
[null,"1,2-Dichlorobenzene","ortho dichlorobenzene",null,null,"N","E","Y","Y","Y","Y","Y","N","N","N","Y","N","2006-05-08","EZ-2 Elite only, or HT series"]
[null,"Dichloroethane",null,null,"DCE","Y","Y","Y","Y","Y","Y","Y","Y","Y","N","Y","Y"]
[null,"Dichloromethane","methylene chloride",null,"DCM","Y","Y","Y","Y","Y","Y","Y","Y","Y","N","Y","Y"]
[null,"Diethyl ether",{"$nulls":3},"N","O","O","N","O","N","O","O","N","N","Y","Y",null,"Must only be used with Inert Gas Purge"]
[null,"Diisopropyl ether","isopropyl ether",null,null,"2007-04-01","Y","Y","Y","Y","Y","Y","Y","Y","N","Y","Y","2009-01-15",null,"RD"]
[null,"1,2-Dimethoxyethane","dimethoxyethane",null,null,"N","O","O","N","O","N","O","O","N","N","Y","N","2011-07-15","Must only be used with Inert Gas Purge","RD"]
[null,"Dimethylacetamide",null,null,"DMAC","-","O","Y","Y","Y","Y","Y","N","N","N","Y","Y",null,"EZ-2 plus / Elite only"]
[null,"Dimethylbenzene","xylene",null,null,"N","O","Y","Y","Y","Y","Y","Y","Y","N","Y","N","2009-10-27","EZ-2 plus / Elite only","RD"]
[null,"Dimethlylformamide",null,null,"DMF","N","O","Y","Y","Y","Y","Y","Y","Y","N","Y","Y",null,"EZ-2 plus / Elite only"]
[null,"Dimethyl sulphoxide",null,null,"DMSO","N","E","Y","Y","Y","Y","Y","N","N","N","Y","Y",null,"EZ-2 Elite only, or HT series"]
[null,"1,4-Dioxane",{"$nulls":3},"2007-04-01","Y","Y","Y","Y","Y","Y","Y","Y","N","Y","Y",null,null,"RD"]
[null,"1,2-Ethanediol","ethylene glycol ","107-21-1",null,"N","N","Y","Y","Y","Y","Y","N","N","N","Y","N","2013-06-14","Very high BP, may need SG temp >40°C"]
[null,"1,2-Ethanedithiol",{"$nulls":3},"N","O","Y","Y","Y","Y","Y","-","-","N","Y","N","2009-03-18","EZ-2 plus / Elite only","RD"]
[null,"2-Ethoxyethanol",{"$nulls":3},"N","O","O","N","O","N","O","O","N","N","Y","N","2011-07-15","Must only be used with Inert Gas Purge","RD"]
[null,"Ethyl acetate",{"$nulls":3},"2007-04-01","Y","Y","Y","Y","Y","Y","Y","Y","N","Y","Y","2007-03-14",null,"RD"]
[null,"Ethyl lactate",{"$nulls":3},"N","O","Y","Y","Y","Y","Y","Y","Y","N","Y","N","2011-09-12","EZ-2 plus / Elite only","RLS"]
[null,"Ethanol",null,null,"EtOH","Y","Y","Y","Y","Y","Y","Y","Y","Y","N","Y","Y"]
[null,"Ethenyl benzene","styrene",null,null,"N","O","Y","Y","Y","Y","Y","Y","Y","N","Y","N","2012-05-02","Best not used in HCl resistant systems due to issues with Hastelloy.  EZ-2 Plus / Elite only","RD"]
[null,"Heptane",{"$nulls":3},"Y","Y","Y","Y","Y","Y","Y","Y","Y","Y","Y","Y","2006-05-19",null,"RD / RLS"]
[null,"Hexane",{"$nulls":3},"Y","Y","Y","Y","Y","Y","Y","Y","Y","Y","Y","Y","2006-04-27",null,"RD"]
[null,"Hydrogen cyanide","Hydrocyanic acid",null,"HCN","N","O","O","N","O","N","O","O","N","Y","Y","N","2011-03-25","Only with Inert Purge.  Note safety issues on handling gas","RD"]
[null,"Hydrogen peroxide",null,null,"H2O2","N","O","N","N","N","N","N","N","N","N","Y","Y","2010-02-17","EZ-2 Plus HCl resistant only - not Elite, not HT series","RD"]
[null,"Limonene",{"$nulls":3},"N","E","Y","Y","Y","Y","Y","N","N","Y","Y","N","2009-02-11","BP too high for miVac, EZ-2 Elite or HT series only","RD"]
[null,"Methanol",null,null,"MeOH","Y","Y","Y","Y","Y","Y","Y","Y","Y","N","Y","Y"]
[null,"Methyl tertiary butyl ether","2-Methoxy-2-methylpropane",null,"MTBE","2007-04-01","Y","Y","Y","Y","Y","Y","Y","Y","N","Y","Y","2007-03-14",null,"RD"]
[null,"Methoxybenzene","anisole",null,null,"N","E","Y","Y","Y","Y","Y","N","N","N","Y","N","2011-05-17","EZ-2 Elite only, or HT series","RD"]
[null,"1-Methoxy-2-propanol","Propylene Glycol Methyl Ether",null,"PGME","N","O","Y","Y","Y","Y","Y","Y","Y","N","Y","N","2011-12-15","EZ-2 plus / Elite only","RD"]
[null,"1-Methoxy-2-propanol acetate","Propylene Glycol Methyl Ether Acetate",null,"PGMEA","N","O","Y","Y","Y","Y","Y","Y","Y","N","Y","N","2011-07-26","EZ-2 plus / Elite only","RD"]
[null,"Methyl aldehyde","Formaldehyde",null,null,"N","Y","Y","Y","Y","Y","Y","Y","Y","N","Y","N","2011-08-19",null,"RD / RLS"]
[null,"N-Methylpyrrolidone",null,null,"NMP","N","E","Y","Y","Y","Y","Y","N","N","N","Y","Y",null,"EZ-2 Elite only, or HT series"]
[null,"Methyl methacrylate",null,null,"MMA","N","*","*","*","*","*","*","*","*","N","Y","N","2011-07-26","MMA is light sensitive and polymerises exothermically when initiated, therefore if the customer has made full risk assesment of this, and has inhibited their solution to prevent dangerous reactions, may be used with regards to the evaporator, however they must control the solution with regards to all aspects of safety.","RD"]
[null,"Naphthalene",{"$nulls":3},"N","E","Y","Y","Y","Y","Y","N","N","N","Y","N","2011-07-26","EZ-2 Elite only, or HT series.  May need high temperature","RD"]
[null,"Octane",{"$nulls":3},"2007-04-01","Y","Y","Y","Y","Y","Y","Y","Y","Y","Y","N","2010-07-05"]
[null,"Octanol",{"$nulls":3},"N","E","Y","Y","Y","Y","Y","N","N","N","Y","N",null,"EZ-2 Elite only, or HT series"]
[null,"Pentane",{"$nulls":3},"N","O","O","N","O","N","O","O","N","Y","Y","Y",null,"Must only be used with Inert Gas Purge"]
[null,"1-Pentanol","Amyl Alcohol",null,null,"-","O","Y","Y","Y","Y","Y","Y","Y","N","Y","Y","2006-06-12","EZ-2 plus / Elite only"]
[null,"Petroleum ether",{"$nulls":3},"2007-04-01","Y","Y","Y","Y","Y","Y","Y","Y","Y","Y","Y","2011-09-19",null,"RD"]
[null,"1,2-Propanediol","Propylene Glycol",null,null,"N","N","Y","Y","Y","Y","Y","N","N","Y","Y","N","2010-07-14",null,"RD"]
[null,"Propan-1-ol",{"$nulls":3},"Y","Y","Y","Y","Y","Y","Y","Y","Y","Y","Y","Y"]
[null,"Propan-2-ol","Iso-Propyl-Alcohol",null,"IPA","Y","Y","Y","Y","Y","Y","Y","Y","Y","Y","Y","Y"]
[null,"Propane-1,2,3-triol","Glycerol / Glycerine",null,null,"N","N","N","N","N","N","N","N","N","N","N","N","2013-09-25","BP 290°C is too high","RD"]
[null,"Propanone","Acetone",null,null,"Y","Y","Y","Y","Y","Y","Y","Y","Y","N","Y","Y","2006-08-17",null,"RD"]
[null,"2-Propyn-1-ol","Propargyl alcohol",null,"HC2CH2OH","N","Y","Y","Y","Y","Y","Y","Y","Y","N","Y","N","2013-07-25","Not in miVac due to plastics compatibility questions","RD"]
[null,"tert-Butyl alchohol","Tertiary butanol",null,"t-BuOH","Y","Y","Y","Y","Y","Y","Y","Y","Y","Y","Y","Y","2009-10-27",null,"RD"]
[null,"Tetraethylene glycol dimethyl ether","Tetraglyme",null,null,"N","N","N","N","N","N","N","N","N","N","N ","N","2013-09-25","BP 275°C is too high","RD"]
[null,"Tetrahydrofuran",null,null,"THF","2007-04-01","Y","Y","Y","Y","Y","Y","Y","Y","N","Y","Y","2007-03-14",null,"RD"]
[null,"1,2,4-Trichlorobenzene",{"$nulls":3},"N","E","Y","Y","Y","Y","Y","N","N","Y","Y","N","2008-04-02","Run at  >50°C to boil.  Must have good vacuum","RD"]
[null,"2,2,2-Trifluoroethanol","trifluoroethyl alcohol",null,"TFE","Y","Y","Y","Y","Y","Y","Y","Y","Y","N","Y","N","2013-09-26","Suggested incompatibilities with nylon","RD"]
[null,"2-Trifluoromethyl-3-ethoxydodecafluorohexane","Novec / HFE 7500",null,null,"N","O","Y","Y","Y","Y","Y","Y","Y","N","Y","N","2013-11-08","EZ-2 plus / Elite only","RD / RLS"]
[null,"1,2,4-trimethylbenzene",{"$nulls":3},"N","E","Y","Y","Y","Y","Y","N","N","N","Y","N","2013-04-10",null,"RD"]
[null,"Trichloroethylene",{"$nulls":3},"N","N","N","N","N","N","N","N","N","N","N","N","2008-04-09","Incompatible with oxidizing agents, aluminium, magnesium, strong bases, reducing agents. Light-sensitive. Reacts violently with many metals, ozone, potassium nitrate, potassium hydroxide, sodium hydroxide.","RD"]
[null,"Triethylammonium acetate",null,null,"TEAA","O","O","Y","Y","Y","Y","Y","Y","Y","N","Y","N","2013-10-03","Quattro Pump / EZ-2 plus / Elite only","RD"]
[null,"Triethylsilane",{"$nulls":3},"N","O","O","N","O","N","O","O","N","N","Y","N","2008-08-21","Inert Purge Mandatory.  Dissociates to evolving Hydrogen in water","RD / RLS"]
[null,"Triisopropylsilane",{"$nulls":3},"N","Y","Y","Y","Y","Y ","Y","N","N","N","Y","N","2008-08-22"]
[null,"2,2,4-Trimethylpentane","isooctane",null,null,"Y","Y","Y","Y","Y","Y","Y","Y","Y","Y","Y","N","2011-07-20",null,"GB"]
[null,"Toluene",{"$nulls":3},"2007-04-01","Y","Y","Y","Y","Y","Y","Y","Y","N","Y","Y","2007-03-14",null,"RD"]
[null,"Water",null,null,"H2O","Y","Y","Y","Y","Y","Y","Y","Y","Y","Y","Y","Y"]